WIM Terminal — System Overview

ADJ Engineering Pvt. Ltd. — Current System (basic_interface_ind570.py)

Covers the current production path only: basic_interface_ind570.py (PyQt5 GUI) reading from a Mettler-Toledo IND570 weighing terminal wired to a Mettler-Toledo load cell (not a raw strain-gauge bridge).

Note: An older USB-packet / Wheatstone-arm architecture (bridge.c, basic_interface.c, gui_interface.py, 45-byte binary packets) previously existed in this repo and has since been removed. basic_interface_ind570.py is the sole application going forward.

## 1. What the system is

A live weight-display application for a weighbridge/WIM (Weigh-In-Motion) setup:

Mettler-Toledo load cell
        │  (analog load signal)
        ▼
   IND570 terminal            ← does ADC, scaling, and outputs a weight stream
        │  Ethernet (TCP), plain ASCII text
        ▼
basic_interface_ind570.py     ← this application (PyQt5 GUI)
        │
        ▼
   On-screen 4-channel display (7-segment style)

The IND570 is the only thing that talks to the load cell directly. The application never touches the sensor signal — it only reads a text stream the IND570 already produces over the network. See ETHERNET_DATA_TRANSMISSION.md for the full detail on that link.

## 2. Why a "4-channel" layout, with one physical load cell

The UI (ChannelColumn, "4x4 Basic Interface") renders four identical channel columns, but the current hardware is a single IND570 + single Mettler-Toledo load cell feed. All four columns read the same underlying value from SharedState — the 4-channel layout is inherited from an earlier 4-load-cell platform design and kept for UI consistency. There is currently no per-channel hardware wiring; it is one live value replicated across four display columns.

## 3. Application architecture

### 3.1 Process / threading model

| Component | Type | Responsibility |
| --- | --- | --- |
| reader_thread() | Background threading.Thread (daemon) | Owns the TCP socket to the IND570; parses incoming lines; pushes values into SharedState |
| SharedState | Thread-safe object (threading.Lock) | Holds the latest averaged result, validity flag, byte counter; read/written from both the reader thread and the UI thread |
| MainWindow (Qt QMainWindow) | Main/UI thread | Builds the window, drives a QTimer at 100 ms to poll SharedState and redraw |

The reader thread and the UI thread never touch each other's data directly — everything passes through SharedState's locked push() / snapshot() methods. This is the standard pattern for keeping a Qt UI responsive while a blocking socket read happens elsewhere.

### 3.2 Startup sequence

main() creates the QApplication, sets the Fusion style, and constructs MainWindow.

MainWindow.__init__ builds all UI panels, creates a SharedState(avg_win=100), sets a threading.Event (_running) to signaled, and starts reader_thread as a daemon thread.

A QTimer fires every 100 ms calling _update_ui(), which is the only place the UI actually reflects new data.

The reader thread independently loops: connect → read → parse → push, retrying every 2 seconds on failure, for as long as _running stays set.

### 3.3 Shutdown

MainWindow.closeEvent clears the _running event, which lets the reader thread's loop exit on its next check; the thread is a daemon so it won't block process exit even if it's mid-recv().

## 4. UI layout

┌───────────────────────────────────────────────────────────────────┐
│  [COM-Port controls: Start | Stop | port dropdown | ...]  (legacy) │
│                                          [Save coefficients]       │
│  Chanel data                                                       │
├───────────┬───────────┬───────────┬───────────────────────────────┤
│ Channel 1 │ Channel 2 │ Channel 3 │ Channel 4                      │
│ [7-seg]   │ [7-seg]   │ [7-seg]   │ [7-seg]      (raw display)     │
│ temp =    │ temp =    │ temp =    │ temp =                         │
│ k = ...   │ k = ...   │ k = ...   │ k = ...      (x2 rows each)     │
│ [7-seg] kg│ [7-seg] kg│ [7-seg] kg│ [7-seg] kg   (value display)   │
│ Cal value │ Cal value │ Cal value │ Cal value                      │
│ [Calibrate][Zero][Cancel zero]  (per column, x4)                  │
├───────────┴───────────┴───────────┴───────────────────────────────┤
│ Average by [100] measurements | Mode: [kg ▾] | Δ Delta | Live      │
│ Base: ...   |   Last packages: — Bytes on port: N   | Exit program │
└───────────────────────────────────────────────────────────────────┘

### 4.1 Top bar — "COM-Port controls"

A QGroupBox with Start / Stop buttons, a port dropdown pre-populated with COM7 and 192.168.0.1:1702, and Update port list / Save port name buttons.

Important: this entire group is cosmetic/legacy. None of these controls read or change the actual connection target — the app always connects to the hardcoded IND570_HOST:IND570_PORT (basic_interface_ind570.py:22-23). Start/Stop do work (they toggle the reader thread via _running), but the COM port dropdown and its buttons are not wired to anything.

### 4.2 Channel columns (×4)

Each ChannelColumn:

Title — "Chanel N"

Top 7-segment display — currently always shows a static "0" (or blank when zeroed); not driven by live data in the current build

temp = — read-only field, always "0" (no temperature input exists in this system)

Two "k =" fields — free-text coefficient inputs, default 0.0067563; present in the UI but not read by any calculation in the current code

Value display (7-segment) + unit label — this is the live figure: shows the averaged/converted weight, updated every 100 ms

Calibration value field + Calibrate button — sets a per-column additive offset (_cal_offset) applied to that column's displayed value only

Zero / Cancel zero buttons — per-column blanking (shows " 0" on both displays), independent of the other three columns and independent of the global tare

### 4.3 Bottom bar

Average by N measurements (QSpinBox, 1–1000) — live-adjusts the rolling-average window size in SharedState

Mode dropdown — kg, 2 × kg, Tons, 2 × Tons; scales the live kg value and changes the unit label (see §5)

Δ Delta — captures current value as a base for all four columns (see §5.1)

Live — clears the delta base, returns to live display

Base: … label — shows the currently captured delta base, blank when not set

Last packages / Bytes on port — status line; shows a red error message if the TCP connection is down

Exit program — closes the window

## 5. Modes and calculations

| Mode | Displayed value | Unit shown |
| --- | --- | --- |
| kg | live averaged kg | kg |
| 2 × kg | live averaged kg × 2 | kg |
| Tons | live averaged kg (no conversion factor applied) | T |
| 2 × Tons | live averaged kg × 2 | T |

Note: Note: "Tons" mode currently displays the same numeric value as kg mode, just with the unit label changed to T — there is no ÷1000 conversion applied in _update_ui or _on_mode_changed. If a true kg→ton conversion is required, this is a code change, not a configuration option.

### 5.1 Delta (Δ) feature

Pressing Δ Delta captures the current mode-adjusted value and the current raw kg (_on_delta), storing both as the "base" for every channel column.

While a base is set, each column's value display shows:

base + (current_raw_kg − base_raw_kg) / 1000

i.e., the base plus a small decimal representing the change since the base was captured — useful for watching small incremental loads relative to a much larger base weight without losing resolution.

Live clears the base on all columns and reverts to showing the live value directly.

### 5.2 Per-column Zero / Cancel zero

Independent of Delta: pressing a column's Zero button blanks that column's two displays to 0; Cancel zero resumes normal display. This does not affect the global average, the other columns, or the calibration offset.

### 5.3 Per-column Calibrate

Typing a number into Calibration value and pressing Calibrate stores it as _cal_offset, added directly to that column's displayed value only (update_value, basic_interface_ind570.py:329-347). It does not affect the underlying averaged kg in SharedState, only what that one column shows.

## 6. Averaging and smoothing

SharedState keeps a collections.deque capped at the configured window size (default 100, adjustable 1–1000 via the spinner).

Every parsed kg value from the IND570 stream is appended; the displayed result is the arithmetic mean of everything currently in the deque.

Changing the spinner value creates a new, empty deque at the new size (set_avg_window) — the average resets and starts refilling immediately, it does not resample the old buffer.

## 7. Error / connection states

_update_ui() checks SharedState.valid:

| valid | Meaning | UI behavior |
| --- | --- | --- |
| 0 | No data yet | Status bar unchanged from initial text |
| 1 | Receiving data normally | Status bar shows byte count in black; channel displays update |
| -1 | Socket exception (connect/read failed) | Status bar shows ERROR: Cannot connect to <host>:<port> in red; channel displays stop updating (last value stays on screen) |

The reader thread retries the connection every 2 seconds indefinitely while _running is set — no manual reconnect action is needed if the IND570 or network comes back.

## 8. Build / packaging

Entry point: basic_interface_ind570.py

Dependency: PyQt5 (see requirements.txt)

Packaging: StrainGauge.spec (PyInstaller) bundles adj_logo_small.png and adj_logo.jpeg alongside the executable, named StrainGauge, windowed (no console)

CI: .github/workflows/build-windows.yml builds StrainGauge.exe on windows-latest on every push to main and on version tags (v*), uploading it as a workflow artifact and attaching it to the GitHub Release for tagged builds

## 9. Known gaps between UI and behavior (for future cleanup)

These are not bugs preventing operation, but UI elements that currently do nothing or less than they visually imply — worth knowing so they aren't mistaken for configuration options:

COM-port dropdown, Update port list, Save port name — not wired to the connection (see §4.1)

Save coefficients button (top-right) — not wired to anything

Top 7-segment display per column, temp = field, both k = fields — not driven by any live calculation

Tons / 2 × Tons modes — do not apply a kg→ton scale factor, only relabel the unit

If any of these are meant to become functional (e.g., configurable IND570 host/port from the dropdown, real per-arm coefficients, true ton conversion), that's a scoped follow-up — flag which ones matter and it can be planned separately.