OIML R 106-1 Type Evaluation Prep

Indicator Module — ADJ Engineering Pvt. Ltd.

Scope: this covers only the electronics/software module we're responsible for —

IND570 + `basic_interface_ind570.py` — submitted to the lab as a *module*

under OIML R 106-1 §5.1.4, not the complete rail-weighbridge (load receptor,

rails, aprons, wagon recognition are handled elsewhere and tested separately).

The deployed field system does dynamic in-motion weighing; what we've built

here is the display/indicator layer, which is why it's evaluated as a module

rather than a complete instrument.

## 1. Why "module" evaluation (§5.1.4)

Testing a complete rail-weighbridge as one unit isn't practical here since

we don't own the mechanical/rail portion. §5.1.4 explicitly allows this:

Note: "the manufacturer may define and submit modules to be examined separately...

Note: where modules are manufactured and/or placed on the market as separate units

Note: to be incorporated in a complete instrument"

Typical modules listed in §0.2.6 include: *load cell, indicator, analogue or

digital data processing device, weighing module, terminal, primary display.*

Our module maps to "indicator" (§0.2.6.4) — displays the weighing result,

does not itself have the load receptor.

### 1.1 Error apportionment (§5.1.4.1)

Each module gets a fraction pi of the complete instrument's mpe. For an

"electronic indicator" specifically, Table 5 (§5.1.4.1) gives:

| Performance criteria | Electronic indicator pi |
| --- | --- |
| Combined effect (non-linearity, hysteresis, temp effect on span, repeatability) | 0.5 |
| Temperature effect on no-load indication | 0.5 |
| Power supply variation | 1 |
| Span stability | 1 |

These fractions apply once the accuracy class of the *complete* instrument is

known — confirm with whoever owns the mechanical/rail module what accuracy

class (0.2 / 0.5 / 1 / 2) the complete system targets, since our module's

error budget is a fraction of that, not a number we can pick independently.

Action needed: get the target accuracy class from the mechanical/rail

team before the lab can compute pass/fail thresholds for our module's tests.

## 2. Documentation to submit (§5.1.1)

Checklist, with status against what exists in this repo today:

| Item | Status | Notes |
| --- | --- | --- |
| Metrological characteristics of the module | ⬜ Not written | Need scale interval (d) as configured/expected from IND570, resolution |
| Functional description of components | ✅ SYSTEM_OVERVIEW.md | Covers threading model, UI, data flow — usable as a base |
| Drawings/diagrams/photos | ⬜ Not prepared | A block diagram of IND570 → TCP → app → display would satisfy this |
| Interface documentation (§4.3.5.1) | ✅ ETHERNET_DATA_TRANSMISSION.md | Covers the TCP link; needs light editing into a formal "list of commands" format (see §5 below) |
| Software documentation (§3.8.1) | 🟡 Partial | See §3 below — several sub-items missing |
| Software identification (§0.2.8.5, §3.11.1) | ✅ Done this session | SOFTWARE_ID = "1.0.0" constant, shown in window title |
| Description of securing components/interlocks | ⬜ Not applicable / not implemented | See §4 — flagging as a gap, not a false "N/A" |
| Data storage device description (§3.5) | ⬜ N/A currently | App does not persist measurement records — flag to lab as "not implemented in this module" rather than omitting |
| Operating manual | 🟡 Partial | SYSTEM_OVERVIEW.md functions as one but isn't labeled/structured as such yet |

## 3. Software documentation gaps (§3.8.1)

§3.8.1 lists what manufacturer software documentation "may include":

a) Description of legally relevant software — ⬜ not done. Need to state

plainly: the entire basic_interface_ind570.py is legally relevant per

§0.2.8.6 ("if no software separation exists, the whole software is to be

considered as legally relevant") — there is currently no separation

between metrologically relevant code (parsing, averaging, calibration)

and non-relevant code (UI chrome). Simplest honest answer for the lab:

declare the whole module as legally relevant rather than attempt a

separation that doesn't exist yet.

b) Description of the accuracy of the measuring algorithms — ⬜ not

written. This should describe: the rolling-average method (SharedState,

arithmetic mean over a configurable window), the unit-normalization

conversions (kg/g/lb/t → kg), and explicitly flag that **"Tons" mode

currently does not apply a ÷1000 conversion** (see SYSTEM_OVERVIEW.md

§5) — this is exactly the kind of accuracy claim a type evaluator will

check, so it needs to either be fixed or explicitly documented as

intentional before submission.

c) Description of the user interface, menus and dialogues — ✅ covered

in SYSTEM_OVERVIEW.md §4.

d) Unambiguous software identification — ✅ done this session.

e) Description of the embedded software — N/A (this runs on a general

PC, not embedded firmware) — should be stated as such rather than left

blank.

f) Overview of system hardware — 🟡 partial. ETHERNET_DATA_TRANSMISSION.md

covers the network topology; a one-paragraph "runs on a standard Windows

PC, no embedded hardware" statement would close this out.

g) Means of securing software — ⬜ real gap, see §4.

h) Operating manual — 🟡 as above.

## 3a. Descriptive markings owed by this module (§3.11)

§3.11 requires markings "at each location having a mass indicating device" —

our display screen is that location. §3.11.4 explicitly allows markings to

be shown on a software-controlled display instead of a physical plate,

provided:

Max, Min, and d are displayed as long as the instrument is switched on

other markings may be shown on manual command

this is described in the type approval certificate

the markings are treated as device-specific parameters (§0.2.8.4) — i.e.

subject to the same securing requirements as §3.8/§3.9

Current UI (basic_interface_ind570.py) does not display any of these

persistently. Concretely missing from the always-visible screen:

| Marking | Required by | Currently shown? |
| --- | --- | --- |
| Max (maximum capacity) | §3.11.4 (always-on) | ❌ no |
| Min (minimum capacity) | §3.11.4 (always-on) | ❌ no |
| d (scale interval) | §3.11.4 (always-on) | ❌ no |
| Software identification | §3.11.1 | ✅ now in window title (SOFTWARE_ID) |
| Manufacturer name/ID | §3.11.4 (on the plate even when display-controlled) | ❌ no |
| Supply voltage | §3.11.4 | N/A — not applicable to PC software; state explicitly |
| Type approval sign | §3.11.4 | Not yet issued — placeholder until approval granted |
| Accuracy class | §3.11.2.1 (code, on manual command acceptable) | ❌ no |
| Weighing method | §3.11.1 | ❌ no — should state "static display of IND570-sourced value" |

This is a real, fixable UI gap — if the plan is to rely on §3.11.4's

allowance for a software-controlled display rather than a physical plate,

the app needs a persistent status area (or footer) showing at minimum

Max/Min/d, and a way to bring up the rest "on manual command" (e.g. an

About/Info dialog). Worth deciding whether to build this before submission

or handle descriptive markings via a physical plate on the IND570 instead,

which may already satisfy this if the IND570 itself carries these markings.

## 4. Securing requirements (§3.8.2, §3.9) — genuine gaps

This is the area with the most actual work, not just documentation:

§3.8.2(a): "legally relevant software shall be adequately protected

against accidental or intentional changes." Currently: the .py source

and the packaged .exe have no integrity protection, no code signing, no

checksum verification at runtime. Worth discussing with the lab what's

acceptable for a module submission — a common minimum is documenting the

build/release process (this repo's GitHub Actions workflow) as the

control, plus the software-identification string added this session as

the means of detecting drift.

§3.9.2(a): access control by code/key for anything that changes

legally-relevant parameters. Currently: no such control exists —

IND570_HOST/IND570_PORT and the averaging window are either hardcoded

or freely editable via the UI (the "Average by N measurements" spinner,

calibration offset fields) with no authorization step. If the calibration

offset (_cal_offset, per-channel) is considered a legally-relevant

parameter, this is a real compliance gap, not just a documentation one —

flag this specifically to whoever is coordinating the submission, since it

may need a code change (e.g., a password-gated calibration mode) before

the lab will accept it.

§3.9.2(b): intervention records (who changed what, when) — not

implemented. No audit trail exists in the app.

**These three are the items most likely to actually block type approval as

currently built** — worth raising early with the lab/certifying body rather

than discovering it during evaluation.

## 5. Interface documentation format (§4.3.5.1)

The lab wants, formally: (a) list of all commands, (b) description of the

software interface, (c) list of all commands together [sic — likely

duplicate wording in the standard], (d) effect of each on functions/data.

ETHERNET_DATA_TRANSMISSION.md already has the substance (TCP client,

port 1702, ASCII line protocol, regex-based parsing) but not in a

command-list format, because the IND570 link isn't a command/response

protocol — it's a continuous one-way data stream. Recommend explicitly

stating this asymmetry to the lab: "the interface accepts no commands from

this module; it is a read-only continuous data stream from the IND570,"

which satisfies (a)–(d) by stating there are no commands rather than

leaving the section blank.

## 6. What the lab will test on our module specifically (Annex A, as a module)

Per §5.1.2, "Instruments may be tested on the premises of the metrological

authority" — for a module, expect at minimum:

A.6.2 Agreement between indicating and printing devices — N/A, no

printing device exists in this module; state explicitly.

A.6.4 Functionality at voltages below minimum operating voltage — N/A

for software running on a general PC; the PC's own power behavior isn't

ours to certify. Flag this boundary explicitly to the lab.

A.7.2 / A.7.3 Influence factor and disturbance tests (temperature,

humidity, AC voltage, ESD, bursts, surges, EMC) — these typically apply to

the indicator hardware (the IND570 itself, which Mettler-Toledo has

presumably already type-approved separately) rather than to display

software running on a PC. Worth clarifying with the lab whether our

module even needs Annex A.7 testing, or whether that burden sits with the

IND570's own existing OIML approval and only the *software's calculation

accuracy* (A.5, static weighing performance) needs re-verification here.

A.8 Span stability — depends on whether "span" means anything at the

software layer (it doesn't add/remove span, only displays/averages what

IND570 sends) — likely N/A, but confirm rather than assume.

**Recommend getting this scoping confirmed by the lab in writing before

investing further prep time** — a large fraction of Annex A is written for

instruments with their own load cell/ADC/environmental exposure, which this

software module does not have.

## 7. Immediate action items

Confirm with the lab/certifying body: does OIML R 106-1 module evaluation

even apply the way we're assuming, or is there a more specific pathway

for "display software only, IND570 already certified separately"?

Get the target accuracy class from the mechanical/rail team (needed for

§5.1.4.1 error apportionment).

Decide whether to fix the "Tons mode doesn't actually convert" issue

(SYSTEM_OVERVIEW.md §5) before submission, or document it as

intentional/known — a type evaluator will find this either way.

Decide how to handle §3.9.2(a) access control on calibration/config

parameters — likely needs a code change, not just documentation.

Confirm with the lab whether Annex A.7 influence/disturbance testing

applies to this software module or sits entirely with the IND570's own

prior approval.